Safety Precautions

1.  Do not leave an Electrothermal IA9100 on if you are not actually using it

2.  Do not leave the Electrothermal IA9100 running unattended.

General Information

1.  Use no more than 1 mm of sample in the capillary tube.  A larger sample may lead to non-uniform heating, and poor results.

2.  It is virtually impossible to determine precisely a melting point (i. e. 123 oC).  Rather, chemists actually determine a melting range (i. e. 122 – 124 oC).  The first temperature in the melting range is recorded when the first signs of liquid droplets appear in the sample, and the second temperature is recorded when the entire sample becomes liquid.

3.  Sometimes, the following may be observed:

i.  The sample will “settle” in the melting point tube

ii.  Solvent will condense on the cooler portions of the tube

It is important to note that these phenomena are not melting.

I.  Establishing a Precise Melting Point

Once an approximate melting point has been established using an Mel-Temp or similar piece of equipment (or the melting point of the substance is known from the literature), the Electrothermal IA9100 apparatus can be used to find the precise melting range for the substance.  The Electrothermal IA9100 automatically regulates the rate of temperature rise to be 1 oC/min – which is the optimum value for determining precise melting points.

1.  Turn on the apparatus using the on/off switch on the back of the power supply.  If the power supply is already on, press the CLEAR button  at least 4 times.

2.  Enter the starting temperature (the “set point” – usu. 5 oC below the measured approximate melting point or literature value).

NOTE:  The arrowed UP key  raises the set point in increments of 10 oC, while the arrowed DOWN key  lowers the set point in increments of 1 oC.  Use the appropriate combination of these keys to set the starting temperature.  For example, to enter a set point of 107 oC, press the UP key  11 times followed by pressing the DOWN key  3 times:

3.  Go to the set point by pressing the GOTO key .  When the set point temperature is reached, the left hand L.E.D. on the GOTO key will illuminate  and three “beeps” will sound to indicate that the temperature has stabilized.

4.  Insert a melting point capillary tube containing a sample into the holder.

5.  Initate the ramp (1 oC/min) by pressing the GOTO button.  The left hand L.E.D. will go out and the right hand L.E.D. will illuminate .

6.  Observe the precise melting range of the sample by looking through the observation window.

7.  When finished, press the CLEAR button  at least 4 times or turn the unit off.

# Text boxes

GOTO key

UP and DOWN keys (for establishing set point)

CLEAR key

Observation window