Safety Precautions

1.  Never insert a room temperature thermometer into a hot Mel-Temp.  It may shatter.

2.  Do not leave a Mel-Temp on if you are not actually using it, or leave it running unattended.

3.  Do not heat the Mel-Temp beyond the upper temperature of the thermometer.

4.  Do not touch the heating block – it gets very HOT!

General Information

1.  Use no more than 2 mm of sample in the capillary tube.  A larger sample may lead to non-uniform heating, and poor results.

2.  It is virtually impossible to determine precisely a melting point (i. e. 123 oC).  Rather, chemists actually determine a melting range (i. e. 122 – 124 oC).  The first temperature in the melting range is recorded when the first signs of liquid droplets appear in the sample, and the second temperature is recorded when the entire sample becomes liquid.

3.  Sometimes, the following may be observed:

i.  The sample will “settle” in the melting point tube

ii.  Solvent will condense on the cooler portions of the tube

It is important to note that these phenomena are not melting.

I.  Establishing an Approximate Melting Point

The Mel-Temp apparatus is the most efficient way to quickly establish an approximate melting point for substances whose melting point is unknown.  An approximate melting point can be taken using a rate of temperature rise of about 7 – 10 oC/min.

1  Make sure a thermometer is inserted into the Mel-Temp.  If not, place the bulb end of a thermometer (with an upper temperature of at least 250 oC) in the opening.

2.  Turn on the apparatus using the on/off switch

3.  Insert melting point capillary tube into the holder

4.  Turn the voltage control knob to achieve a rate of temperature rise of about 7 – 10 oC/min.

5.  Observe the approximate melting range of the sample by looking through the observation window.  If the rate of temperature rise slows to less than 5 oC/min, turn the voltage control knob to a higher value to increase the rate.

6.  When finished, turn the voltage control knob to “0” or turn the unit off.

II.  Establishing a Precise Melting Point

Once an approximate melting point has been established (or if the melting point of the substance is known from the literature) the Mel-Temp apparatus can be used to find the precise melting range for the substance.  To establish a precise melting point, a rate of temperature rise of 1 – 2 oC/min must be used.

1  Make sure a thermometer is inserted into the Mel-Temp.  If not, place the bulb end of a thermometer (with an upper temperature of at least 250 oC) in the opening.

2.  Turn on the apparatus using the on/off switch

3.  Turn the voltage control knob so as to establish a starting temperature about 5 oC below the approximate melting point found from part I (or about 5 oC below the literature value for the substance)

4.  Insert melting point capillary tube into the holder.

5.  Adjust the voltage control knob so as to establish the rate of temperature rise to be 
1 – 2 oC/min

6.  Observe the precise melting range of the sample by looking through the observation window.  If the rate of temperature rise slows to less than 0.3 oC/min, turn the voltage control knob to a higher value to increase the rate.

7.  When finished, turn the voltage control knob to “0” or turn the unit off.

# Text boxes

http://spot.pcc.edu/~gbackes/CH222/Labs222.W05/Exp.12.Synthesis.Aspirin_files/image016.jpg

Mel-Temp Apparatus

Heating block